## Slide 1: Introduction to Pollynet_Processing_Chain (Picasso)
Zhenping Yin
2019-07-21

## Slide 2: Outline
Download and Setup
Hands-on to support new pollyxt

## Slide 3: PollyNET Workflow
PollyNET processing program
PollyNET Measurements
PollyNET Data Server

## Slide 4: Requirements for Picasso
Interpreter
MATLAB 2014a
Python 3
Python packages: matplotlib, numpy and scipy
Folder structure
Root Dir + pollyType + ‘data_zip’ + yyyymm
Housekeeping data file
Compressed polly data file
Version control

## Slide 5: Download
git clone https://github.com/ZPYin/Pollynet_Processing_Chain
This command will download the code from the remote GitHub repository
One step further…
If you want to develop new features, we recommend you to pull the ‘dev’ branch as well:
git fetch origin dev:dev
(Detailed information about how to play with `git`, please go to ../doc/git_workflow.pptx).
(There is also ‘download’ option under this link)

## Slide 6: Structure of Picasso
Picasso
figures
results
calibration
log
Unzip polly data
todo_filelist
pollynet_places_history_new.txt
pollynet_processing_chain_link.txt
Load the polly config
pollyxt_xxx defaults
pollyxt_xxx module
Single task

## Slide 7: Setup
Picasso Global Config
1. setup the global config file from the template
2. edit the global config file
directories: according to your own requirements. -
to control whether to delete the unzipped polly data after each processing.
python 3 folder: Picasso uses python3 for data visualization if ‘visualizationMode’was set ‘python’
Detailed information can be found in

## Slide 8: Setup
Picasso adds campaign:
Normally, this file was generated from the database. But if there is no connection to the database, you need to add new campaign info manually.
(Be careful about the character of the ‘space’, actually it consists of ‘tab’. So better way is to copy the line before and edit it according to your campaign.)

## Slide 9: Setup
Picasso adds polly config
Copy the ‘template_config.json’ to a new polly config file with the naming of ‘{polly-type}_config_{date}.json’
You may need to change the settings of:
gdas1Site
AERONETSite
depol_cal_ang_p_time
depol_cal_ang_n_time
meteorDataSource
IWV_instrument
intNProfiles
minIntNProfiles
radiosondeFolder
logbookFile
You can find detailed info about the polly config in /doc/polly_config.md

## Slide 10: Setup
Picasso adds the link between polly config file and polly data
config history file works as a lookup table for Picasso to search the corresponding polly config file.
better way to do is to copy the line before and then edit it according to your demands.
Now, everything is really

## Slide 11: Bash Scripts for automating it

## Slide 12: Add a new campaign
pollynet_places_history_new.txt
pollynet_processing_chain_link.txt
Load the polly config
pollyxt_xxx defaults
pollyxt_xxx module
1. New campaign entry
3. New link entry
2. New polly config

## Slide 13: Hands on

## Slide 14: Add a new Polly
pollynet_places_history_new.txt
pollynet_processing_chain_link.txt
Load the polly config
pollyxt_xxx defaults
pollyxt_xxx module
1. New campaign entry
5. New link entry
4. New polly config
3. New polly defaults
2. New polly modules

## Slide 15: Hands on

## Slide 16: Picasso Development
Requirements:
Be familiar with the whole structure and algorithms
Be professional with MATLAB
TODO
Full support of data visualization with MATLAB, which can remove the dependence on Python
Uncertainty analysis
Products output (Level 1 (signal correction), Level 2 (calibration results, extensive and intensive properties))
Layer detection algorithm (aerosol and clouds)
Dual-FOV method for retrieving cloud microphysical properties

## Slide 17: Picasso Structure
Automatic calibration
  lidar calibration
  water vapor calibration
  depolarization calibration
Integrated conventional retrieving algorithms
Extensible programming structure

## Slide 18: Version Control

## Slide 19: Backup

## Slide 20: Rayleigh Fit
Range Corrected Signal:
Attenuated molecular backscatter:
In aerosol-free or nearly aerosol-free (quantitatively?) area, assuming ,
Therefore
Main task is to search the most linear correlated area

## Slide 21: Rayleigh Fit
Normalize the measured signal with attenuated molecular backscatter
Test 1: Pure Rayleigh Conditions
Test 3: White-Noise criterion
Test 4: SNR constrain
Test 2: Near and far range cross criteria
if no other absorptive aerosol layers above it
(Baars et al, AMT, 2016)

## Slide 22: Rayleigh Fit
Normalize the measured signal with attenuated molecular backscatter
Test 1: Pure Rayleigh Conditions
Test 2: Near and far range cross criteria
Test 3: White-Noise criterion
Test 4: SNR constrain
Smaller the mean of the residual is and more Gauss distribution alike the residual is, more possible it could be aerosol-free.

## Slide 23: Shortage of sliding fixed window
Omission: This possible interval will not be found if a large fixed window (like 1 km) was used, since the width of aerosol-free area is only 450m.
More computation: all the sub-areas with a fixed width will be tested.

## Slide 24: Rayleigh Fit – area decomposition
Interested area:
Not intersected
Nearly aerosol-free
Which means the slope of the interested area should be 0 and the shape of interested area should be nearly linear

## Slide 25: Rayleigh Fit - Douglas-Peucker Algorithm
Gifs to show how Douglas- Peucker Algorithm works
Each interval is not intersected.
Each interval is nearly linear, but slope of each interval is still varied.

## Slide 26: Rayleigh Fit
Douglas-Peucker Algorithm
[[z_min1, z_max1], [z_min2, z_max2], … [z_mink, z_maxk]]
Test 1: Slope check
Test 2: Near and far range cross criteria
Test 3: White-Noise criterion
Test 4: SNR constrain
Select the best fit ref interval:
(mean value, standard deviation and slope of the residual, mean square root error from a linear regression, test statistic of Anderson-Darling test)